3000 Series 
Operation Manual

Copyright

Copyright © 1995 by this company.  All rights reserved.  No part of this publication may be reproduced in any form or by any means without the written permission of this company.

Disclaimer

This company makes no representations or warranties, either expressed or implied, with respect to the contents hereof and specifically disclaims any warranties, merchantability or fitness for any particular purpose.  Further, this company reserves the right to revise this publication and to make changes from time to time in the contents hereof without obligation of this company to notify any person of such revision or changes.

Table of Contents

1.	Overview	1

1.1	Introduction	1

1.2	Unpacking and Checking	1

2.	Front, Rear and Bottom Panels	2

2.1	Front Panels	2

2.2	Rear Panels and Bottom Panels	5

3.	Operation	7

3.1	Pre-installation	7

3.2	Setting the Output Voltage and Output Current	7

3.3	Constant Voltage/Constant Current Characteristics	8

3.4	Operation Modes	8

3.4.1	Serial	8

3.4.2	Parallel	9

4.	Maintenance	11

4.1	Changing the Fuse	11

4.2	Changing the Voltage	11

5.	Specifications	12

# 1.	Overview

This 3000 Series Direct Current (DC) Power Supply Operation Manual contains an introduction to the power supply, a description of its functions, the operation procedure, the scope of application and the specifications.  The models are the 3203A, 3203AR, 3203D, 3203DR, 3302A, 3302AR, 3302D, 3302DR, 3303A, 3303AR, 3303D, 3303DR, 3601A, 3601AR, 3601D, 3601DR, 3306D, 3306DR, 3603D, 3603DR, 33010D and 33010DR.

## 1.1	Introduction

The 3000 Series is a single-channel-output DC power supply with the following features:

·	Constant voltage and constant current output

·	Short circuit and  external input protection

·	Allows serial or parallel connection with the same power supply model

### Special Functions on Remote Models

The following functions are available only on the remote models 3203AR, 3203DR, 3302AR, 3302DR, 3303AR, 3303DR, 3601AR, 3601DR, 3306DR, 3603DR and 33010DR:

·	Output voltage readback

·	Output current readback

·	Remote or manual-controlled ON/OFF without using the power switch

·	Control of output voltage and current using external voltage and resistor

·	Remote sensing

## 1.2	Unpacking and Checking

Unpack and check that you have the following items:

·	One power cable

·	This Operation Manual

·	Remote-Control Operation Manual (remote models only)

·	One ACS-002 banana clip

·	Two phone jacks (remote models only)

·	Two 9-pin RS-232C connectors (remote models only)

# 2.	Front, Rear and Bottom Panels

## 2.1	Front Panels

Front panel with digital display for models 3306DR, 3603DR and 33010DR

Figure 1	Front Panels

(1)	Power Switch: Power ON/OFF switch (1=ON, 0 = OFF)

(2)	Output ON/OFF LED:  Red LED for indicating output ON or OFF

(3)	Output ON/OFF Switch: Output ON/OFF control switch (only available in remote models)

(4)	-S Remote Sensing Input Terminal:  For sensing output voltage at the - load terminal (only available in remote models)

(5)	Output Current Adjustment Knob:  For adjusting output current  at constant current (C.C. Mode)

(6)	- Output Terminal: Negative output terminal (black)

(7)	+ Output Terminal: Positive output terminal (red)

(8)	Voltage Indicator

a)	Analog: As indicated by the pointer on the 40x40mm (Class 2.5) meter

b)	Digital:  Displayed in full 3-digit red 0.52" LED

(9)	C.V. Mode LED: Green LED to indicate constant voltage

(10)	C.C. Mode LED: Red LED to indicate constant current

(11)	Current Indicator

a)	Analog: As indicated by the pointer on the 40x40mm (Class 2.5) meter

b)	Digital:  Displayed in full 3-digit red 0.52" LED

(12)	Coarse Output Voltage Adjustment Knob:  For adjusting output voltage at constant voltage  (C.V. mode)

(13)	GND Terminal : Ground terminal (green)

(14) 	+S Remote Sensing Input Terminal:  For sensing output voltage at the + load terminal (only available in remote models)

(15) 	Fine Output Voltage Adjustment Knob:  For fine adjustment of output voltage at constant voltage  (C.V. mode)

(16)	Warning  Label:  The bare parts of the output terminals and fittings are electrical sensing parts.  Do not touch these parts during use.

## 2.2	Rear Panels and Bottom Panels

Rear and bottom panels for models 3203AR, 3203DR, 3302AR, 3203DR, 3303AR, 3303DR, 3601AR and 3601DR

Rear and bottom panels for models 3306DR, 3603DR and 33010DR

Figure 2   Rear and Bottom Panels

(17)	D-Type 9-pin connector (RC1) for Remote Control

(18)	D-Type 9-pin connector (RC2) for Remote Control

NOTE:	Items 17 and 18 are available only on the remote models.  Refer to the Remote-Control Operation Manual for details.

(19)	Power Input Socket

(20)	Input Power Fuse Holder and Input Voltage Selector.  Acceptable input voltage ratings are 100V, 120V, 220V and 240V.  The selected input voltage is set to the position above the   mark (the rear panel above shows that the input voltage is set to 100V).

(21)	Remote Control Function Switches:   Only available on remote models.  Refer to the Remote-Control Operation Manual for details.

# 3.	Operation

## 3.1	Pre-installation

Before you turn on the power:

·	Check that the input voltage from your power source conforms to the voltage rating selected.  Refer  to  Input Voltage Selector in section 2.2.  The tolerance is +10% of the indicated voltage at 50/60 Hz.

·	Place this power supply in a well-ventillated area and do not block the ventillation holes.   Poor heat dissipation leads to overheating which may cause unstable operation and shorten the service life of this equipment.

·	Be sure to use the correct fuse for your model.  Refer to Table 1 for the type of  fuse you should use for a given voltage.

## 3.2	Setting the Output Voltage and Output Current

To set the output voltage and current, follow these steps:

1.	Check that the total load to be connected does not exceed the maximum output voltage and current of this power supply.

2.	Open the circuit between the + and the - output terminals.  Turn the voltage adjustment knob clockwise to the desired output voltage rating.

3.	Turn the current adjustment knob counterclockwise to get the minimum value.

4.	Short the circuit between the + and the - output terminals.  Note that the current rating of the shorting wire should be greater than or equal to the required current.

5.	Turn the current adjustment knob clockwise until the current indicator on the front panel displays the required current rating.

6.	Remove the shorting wire from the + and  the - output terminals.  The power supply returns to the constant voltage mode and is ready to use.

## 3.3	Constant Voltage/Constant Current Characteristics

This power supply operates automatically between constant voltage (C. V. mode) and constant current (C. C. mode) by responding quickly to rapid load changes.  The following figures shows the relative changes between the constant current and the constant voltage modes.

Figure 3   Constant Voltage vs. Constant Current Relation Diagram

## 3.4	Operation Modes

### 3.4.1	Serial

Serial connection allows the connection of two or more power supplies serially to obtain a higher voltage rating (maximum 240V for this power supply).

For serial connection, it is recommended that all power supplies operate under the constant voltage (C.V.) mode.

Figure 4 shows the serial connection between two 3303D power supplies.  One unit is set at 30V/2A and the other at 20V/3A.  Figure 5 shows the serial connection voltage/current output versus load 
changes.

| Figure 4 Serial Connection Diagram | Figure 5 Serial Voltage/Current Output vs. Load Change |  |
| --- | --- | --- |

### 3.4.2	Parallel

Parallel connection only allows thr connection of two or more power supplies of the same model to obtain  a higher current rating (maximum 24A for this power supply).

Figure 6 shows the parallel connection of two 3303D power supplies.  One unit is set at 30V/3A and the other at 20V/2.5A.  Figure 7 shows the parallel connection voltage/current output versus load 
changes.

| Figure 6 Parallel Connection Diagram |  | Figure 7 Parallel Voltage/Current Output vs. Load Change |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |

NOTE:	The difference between the voltage settings for two or more power supplies of the same model connected in parallel should not exceed 15V.  The minimum voltage setting should not be lower than 10V.  If it is necessary to be lower than 10V, the voltage difference should be less than 2V.  The closer to 0V the smaller the difference required.

# 4.	Maintenance

## 4.1	Changing the Fuse

The fuse is located inside the input power fuse holder (refer to Figure 2).   You need to change the fuse when:

·	the fuse is blown out

·	you change the voltage rating

In any case, replace the fuse with one of the same rating.  Refer to Table 1 for the type of fuse used for each model and voltage.

NOTE:	Unplug the power cord before you change the fuse.

## 4.2	Changing the Voltage

To change the voltage, follow these steps:

1.	Use a flathead screwdriver to detach the Input Power Fuse Holder and Input Voltage Selector  unit (refer to Figure 2).

2.	Turn the unit so that the desired input voltage is positioned above the  mark.

4.	Check that the fuse to conform to this new voltage rating.   Refer to Table 1 below for the correct fuse rating.

5.	Replace the Input Power Fuse Holder and Input Voltage Selector unit.

Table 1	Power Supply Weight , Dimensions and Fuse Specification

| Model | Weight (Approx.) |  |  | Dimension W x H x D (mm) |  |  | Fuse Time-Delay Type 6 x 30 mm |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Net | Gross | Machine |  | Package | 100V |  | 120V | 220V | 240V |
| 3203A/AR/D/DR | 4.3 KG | 5.5 KG | 120x160x200 |  | 239x265x394 | 2A |  | 2A | 1A | 1A |
| 3302A/AR/D/DR | 4.3 KG | 5.5 KG | 120x160x200 |  | 239x265x394 | 2A |  | 2A | 1A | 1A |
| 3303A/AR/D/DR | 4.3 KG | 5.5 KG | 120x160x200 |  | 239x265x394 | 2A |  | 2A | 1A | 1A |
| 3601A/AR/D/DR | 4.3 KG | 5.5 KG | 120x160x200 |  | 239x265x394 | 2A |  | 2A | 1A | 1A |
| 3306D/DR | 8.6 KG | 10.1 KG | 230x160x324 |  | 364x265x422 | 4A |  | 4A | 2A | 2A |
| 3603D/DR | 8.6 KG | 10.1 KG | 230x160x324 |  | 364x265x422 | 4A |  | 4A | 2A | 2A |
| 33010D/DR | 13 KG | 14.4 KG | 230x160x366 |  | 364x265x464 | 8A |  | 8A | 4A | 4A |

# 5.	Specifications

| Model Specification | 3203A 3203D | 3302A 3302D | 3303A 3303D | 3601A 3601D | 3306D |  | 3603D | 33010D |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Output Voltage | 0-20V | 0-30V | 0-30V | 0-60V | 0-30V |  | 0-60V | 0-30V |  |
| Output Current | 0-3A | 0-2A | 0-3A | 0-1A | 0-6A |  | 0-3A | 0-10A |  |
| CONSTANT VOLTAGE CHARACTERISTICS |  |  |  |  |  |  |  |  |  |
| Load Regulation | ± 0.01% +2mV |  |  |  |  |  |  |  |  |
| Line Regulation | ± 0.01% +2mV |  |  |  |  |  |  |  |  |
| Ripple & Noise (≦100W) | ≦0.5mVrms |  |  |  |  |  |  |  |  |
| Ripple & Noise (≧100W) | ≦1mVrms |  |  |  |  |  |  |  |  |
| CONSTANT CURRENT CHARACTERISTICS |  |  |  |  |  |  |  |  |  |
| Load Regulation (≦100W) | ≦10mA |  |  |  |  |  |  |  |  |
| Load Regulation (≧100W) | ≦15mA |  |  |  |  |  |  |  |  |
| Line Regulation | ± 0.01% +2mA |  |  |  |  |  |  |  |  |
| Ripple & Noise (≦100W) | ≦1mArms |  |  |  |  |  |  |  |  |
| Ripple & Noise (≧100W) | ≦3mArms |  |  |  |  |  |  |  |  |
| DISPLAY ACCURACY |  |  |  |  |  |  |  |  |  |
| Analog Display (A) Digital Display (D) | Full-scale 3% ≦0.1% +2d |  |  |  |  |  |  |  |  |
| PROGRAMMING SPEED |  |  |  |  |  |  |  |  |  |
| Rise Time (No Load) (Load) | ≦100mS ≦200mS (<6A), ≦ 500mS (<10A), ≦ 1S (≧10A) |  |  |  |  |  |  |  |  |
| Fall Time (No Load) (Load) | ≦2.5S ≦250mS |  |  |  |  |  |  |  |  |
| OUTPUT IMPEDANCE | < 2mΩ + 2μH |  |  |  |  |  |  |  |  |
| RECOVERY TIME | ≦100μS to within 0.1% of set voltage (50% to 100% load change) |  |  |  |  |  |  |  |  |
| FUNCTIONS |  |  |  |  |  |  |  |  |  |
| Series Connection | Different models can be connected in series (≦240V) |  |  |  |  |  |  |  |  |
| Parallel Connection | Same models can be connected in parallel (≦ 24 A) |  |  |  |  |  |  |  |  |
| Master - Slave | N/A |  |  |  |  |  |  |  |  |
| POWER SOURCE | ACV 100/120/220/240 ±10%, 50/60Hz |  |  |  |  |  |  |  |  |
| ACCESSORIES | ACS-002 X1 |  |  |  |  |  |  |  |  |
| DIMENSIONS W X H X D (mm) | 120 X 160 X 200 |  |  |  |  | 230 X 160 X 324 |  |  | 230 X 160 X 366 |

| Model Specification | 3203AR 3203DR | 3302AR 3302DR | 3303AR 3303DR | 3601AR 3601DR | 3306DR |  | 603DR | 33010DR |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Output Voltage | 0-20V | 0-30V | 0-30V | 0-60V | 0-30V |  | 0-60V | 0-30V |  |
| Output Current | 0-3A | 0-2A | 0-3A | 0-1A | 0-6A |  | 0-3A | 0-10A |  |
| CONSTANT VOLTAGE CHARACTERISTICS |  |  |  |  |  |  |  |  |  |
| Load Regulation | ± 0.01% +2mV |  |  |  |  |  |  |  |  |
| Line Regulation | ± 0.01% +2mV |  |  |  |  |  |  |  |  |
| Ripple & Noise (≦100W) | ≦1mVrms |  |  |  |  |  |  |  |  |
| Ripple & Noise (≧100W) | ≦2mVrms |  |  |  |  |  |  |  |  |
| CONSTANT CURRENT CHARACTERISTICS |  |  |  |  |  |  |  |  |  |
| Load Regulation (≦100W) | ≦10mA |  |  |  |  |  |  |  |  |
| Load Regulation (≧100W) | ≦15mA |  |  |  |  |  |  |  |  |
| Line Regulation | ± 0.01% +2mA |  |  |  |  |  |  |  |  |
| Ripple & Noise (≦100W) | ≦1mArms |  |  |  |  |  |  |  |  |
| Ripple & Noise (≧100W) | ≦3mArms |  |  |  |  |  |  |  |  |
| DISPLAY ACCURACY |  |  |  |  |  |  |  |  |  |
| Analog Display (A) Digital Display (D) | Full-scale 3% ≦ 0.1% +2d |  |  |  |  |  |  |  |  |
| PROGRAMMING SPEED |  |  |  |  |  |  |  |  |  |
| Rise Time (No Load) (Load) | ≦100mS ≦200mS (<6A), ≦ 500mS (<10A), ≦ 1S (≧10A) |  |  |  |  |  |  |  |  |
| Fall Time (No Load) (Load) | ≦2.5S ≦250mS |  |  |  |  |  |  |  |  |
| OUTPUT IMPEDANCE | <2mΩ + 2μH |  |  |  |  |  |  |  |  |
| RECOVERY TIME | ≦100μS to within 0.1% of set voltage (50% to 100% load change) |  |  |  |  |  |  |  |  |
| FUNCTIONS |  |  |  |  |  |  |  |  |  |
| Serial Connection | Different models can be connected in series (≦240V) |  |  |  |  |  |  |  |  |
| Parallel Connection | Same models can be connected in parallel (≦24A) |  |  |  |  |  |  |  |  |
| Master - Slave | Yes |  |  |  |  |  |  |  |  |
| REMOTE CONTROL |  |  |  |  |  |  |  |  |  |
| Voltage Programmable By External DC Voltage By External Resistance | 0 to 10V, Control 0 to voltage setting 0 to 10KΩ, Control 0 to maximum voltage |  |  |  |  |  |  |  |  |
| Current Programmable By External DC Voltage By External Resistance | 0 to 10V, Control 0 to current setting 0 to 10KΩ, Control 0 to maximum current |  |  |  |  |  |  |  |  |
| VOLTAGE/CURRENT READBACK AVAILABLE |  |  |  |  |  |  |  |  |  |
| Voltage Readback | 0 to 10V for full-scale voltage |  |  |  |  |  |  |  |  |
| Current Readback | 0 to 0.3V for full-scale current |  |  |  |  |  |  |  |  |
| OUTPUT ON/OFF CONTROL | YES |  |  |  |  |  |  |  |  |
| REMOTE SENSING | Compensate voltage drop Up to 2VDC |  |  |  |  |  |  |  |  |
| POWER SOURCE | ACV 100/120/220/240 ±10%, 50/60Hz |  |  |  |  |  |  |  |  |
| ACCESSORIES | ACS-002 X1, phone-jack plug X2, 9-pin D-type connector X 2 |  |  |  |  |  |  |  |  |
| DIMENSIONS W X H X D (mm) | 120 X 160 X 200 |  |  |  |  | 230 X 160 X 324 |  |  | 230 X 160 X 366 |