6000 Series 
Operation Manual

Copyright

Copyright © 1995 by this company.  All rights reserved.  No part of this publication may be reproduced in any form or by any means without the written permission of this company.

Disclaimer

This company makes no representations or warranties, either expressed or implied, with respect to the contents hereof and specifically disclaims any warranties, merchantability or fitness for any particular purpose.  Further, this company reserves the right to revise this publication and to make changes from time to time in the contents hereof without obligation of this company to notify any person of such revision or changes.

Table of Contents

1.	Overview	1

1.1	Introduction	1

1.2	Unpacking and Checking	1

2.	Front and Rear Panels	2

2.1	Front Panel	2

2.2	Rear Panel	5

3.	Operation	6

3.1	Pre-installation	6

3.2	Setting the Output Voltage and Output Current	6

3.3	Constant Voltage/Constant Current Characteristics	7

3.4	Operation Modes	7

3.4.1	Independent Mode	7

3.4.2	Serial Tracking Mode	8

3.4.3	Parallel Tracking Mode	9

3.5	Fixed 5V/5A Output	10

3.6	Serial Operation	11

3.7	Parallel Operation	12

4.	Maintenance	13

4.1	Changing the Fuse	13

4.2	Changing the Voltage	13

5.	Specifications	14

# 1.	Overview

This 6000 Series Direct Current (DC) Power Supply Operation Manual contains an introduction to the power supply, a description of its functions, the operation procedure, the scope of applications and its specifications.  The models available under this series are the 6302A, 6302AR, 6302D, 6302DR, 6303A, 6303AR, 6303D, 6303DR, 6306A, 6306AR, 6306D, 6306DR, 6603A, 6603AR, 6603D, and 6603DR.

## 1.1	Introduction

The 6000 Series is a dual-channel-output DC power supply with the following features:

·	Twin power output with tracking function for automatic selection of parallel or serial connection

·	Short-circuit protection against external input while providing constant voltage and constant current

·	Allows serial or parallel connection with the same power supply model

·	5V/5A constant-voltage output

### Special Functions on Remote Models

The following functions are available only on the remote models 6302AR, 6302DR, 6303AR, 6303DR, 6306AR, 6306DR, 6603AR and 6603DR:

·	Control of output voltage and current using external voltage and resistor

·	Output voltage readback

·	Output current readback

·	Random master or slave combination

## 1.2	Unpacking and Checking

Unpack and check that you have the following items:

·	One power cable

·	This Operation Manual

·	Remote-Control Operation Manual (remote models only)

·	Two ACS-002 banana clips

·	Two 9-pin RS-232C connectors (remote models only)

# 2.	Front and Rear Panels

## 2.1	Front Panels

Front panel  with digital display for models 6302D, 6302DR, 6303D, 6303DR, 6306D, 6306DR, 6603D and 6603DR

Front panel  with analog meter for models 6302A, 6302AR, 6303A, 6303AR, 6306A, 6306AR, 6603A and 6603AR

Figure 1     Front Panels

(1)	C.V. Mode LED:  Uses green LED to indicate constant voltage at slave power

(2)	Power Switch: Power ON/OFF switch (1=ON, 0 = OFF)

(3)	C.C. Mode LED: Uses red LED to indicate constant current at slave power

(4)	SLAVE Current Indicator

a)	Analog: As indicated by the pointer on the 40x40mm (Class 2.5) meter

b)	Digital:  Displayed in full 3-digit red 0.52" LED

(5)	SLAVE Voltage Indicator

a)	Analog: As indicated by the pointer in the 40 x 40 mm (Class 2.5) meter

b)	Digital:  Displayed in full 3-digit red 0.52" LED

(6)	MASTER Voltage Indicator

a)	Analog: As indicated by the pointer on the 40 x 40 mm (Class 2.5) meter

b)	Digital:  Displayed in full 3-digit red 0.52" LED

(7)	MASTER Current Indicator

a)	Analog: As indicated by the pointer on the 40 x 40 mm (Class 2.5) meter

b)	Digital:  Displayed in full 3-digit red 0.52" LED

(8)	C.V. Mode LED: Uses green LED to indicate constant voltage at  master power

(9)	C.C. Mode LED: Uses red LED to indicate constant current at  master power

(10)	SLAVE Current Adjustment Knob:  For adjusting output current when slave is at constant current     (C.C.) mode

(11)	Slave Voltage Adjustment Knob:  For adjusting output voltage when slave is at constant voltage  (C.V.) mode

(12)	Slave - Output Terminal: Slave negative output terminal (black)

(13)	Slave GND Terminal: Slave ground terminal (green)

(14)	Slave + Output Terminal: Slave positive output terminal (red)

(15) and (18) TRACKING MODE Switches:  The switch setting label is located between these two switches and you can set the power supply to operate in any of the following modes:

IND:  Independent Mode

SER: Series Mode

PAR: Parallel Mode

(16)	5V- Output Terminal : 5V negative output terminal (black)

(17)	5V+ Output Terminal: 5V positive output terminal (red)

(19)	Master - Output Terminal: Master negative output terminal (black)

(20)	Master GND Terminal: Master ground terminal (green)

(21)	Master + Output Terminal: Master positive output terminal (red)

(22)	MASTER Voltage Adjustment Knob:  For adjusting output voltage when  master  power  is at constant voltage  (C.V.) mode

(23)	MASTER Current Adjustment Knob:  For adjusting output current when  master power  is at constant current (C.C.) mode

(24)	Warning  Label:  The bare parts of the output terminals and fittings are electrical sensing parts.  Do not touch these parts during use.

(25)	OVERLOAD LED: 5V/5A overload indicator LED (red)

## 2.2	Rear Panel

Figure 2     Rear Panel

(26)	Remote Control Function Switches: Only available on remote models.  Refer to the Remote-Control Operation Manual for details.

(27)	D-Type 9-pin connector (RC1) for Remote Control

(28)	D-Type 9-pin connector (RC2) for Remote Control

(29)	Power Input Socket

(30)	Input Power Fuse Holder and  Input Voltage Selector:  Acceptable input voltage ratings are 100V, 120V, 220V and 240V.  The selected input voltage is set to the position above the   mark (the rear panel above shows that the input voltage is set to 100V).

# 3.	Operation

## 3.1	Pre-installation

Before you turn on the power:

·	Check that the input voltage from your power source conforms to the voltage rating selected.  Refer  to the Input Power Fuse Holder and  Input Voltage Selector in section 2.2.  The tolerance is +10%  of the indicated voltage at 50/60 Hz.

·	Place this power supply in a well-ventillated area and do not to block the ventillation holes.   Poor heat dissipation leads to overheating which may cause unstable operation and shorten  the service life of this equipment.

·	Be sure to use the correct fuse for your model.  Refer to Table 1 for the type of  fuse you should use for a given voltage.

## 3.2	Setting the Output Voltage and Output Current

To set the output voltage and current, follow these steps:

1.	Check that the total load to be connected does not exceed the maximum output voltage and current of this power supply.

2.	Open the circuit between the + and the - output terminals.  Turn the voltage adjustment knob clockwise until you get the desired output voltage rating.

3.	Turn the current adjustment knob counterclockwise until you get the minimum value.

4.	Short the circuit between the + and the - output terminals.  Note that the current rating of the shorting wire should be greater than or equal to the required current.

5.	Turn the current adjustment knob clockwise until the current indicator on the front panel displays the required current rating.

6.	Remove the shorting wire from the + and  the - output terminals.  The power supply returns to the constant voltage mode and is ready to use.

## 3.3	Constant Voltage/Constant Current Characteristics

This power supply can automatically operate between constant voltage and constant current by responding quickly to rapid load changes.  The following figures shows the relative changes between the constant current and the constant voltage modes.

Figure 3     Constant Voltage/Constant Current

## 3.4	Operation Modes

### 3.4.1	Independent Mode

The master and the slave power can be used independently to generate voltage and current.  To operate in the independent mode, follow these steps:

1.	Set the TRACKING MODE switches on the front panel to IND.

2.	Turn on the POWER switch.

3.	Set the master output voltage and current as described in section 3.2.

4.	Set the slave output voltage and current as described in section 3.2.

### 3.4.2	Serial Tracking Mode

Under this mode, the - Master Output Terminal is automatically shorted with the + Slave Output Terminal.

This serial output is generated from the + Master Output Terminal and the - Slave Output Terminal.  The output voltage is twice the set master output voltage while the output current is the same as the master setting value.

To operate in the serial mode:

1.	Set the TRACKING MODE switches on the front panel to SER.

2.	Turn on the POWER switch.

3.	Turn the SLAVE Current Adjustment Knob clockwise to maximum.

4.	Set the master output voltage and current as described in section 3.2.

The following figures show two different serial operations:

| The output voltage is twice the set master output voltage. The output terminal is connected from the + Master Output Terminal to the - Slave Output Terminal. |  | The output voltages for the master and the slave are the same. You can apply two independent loads or a load requiring both positiveand negative voltages as long as the positive and negative voltages are the same. |
| --- | --- | --- |

Figure 4     Serial Connections

### 3.4.3	Parallel Tracking Mode

Under this mode, the - Master Output Terminal is shorted with the - Slave Output Terminal while the + Master Output Terminal is shorted with the + Slave Output Terminal automatically.

This parallel output is generated from the + Master Output Terminal and the -Master Output Terminal (or - Slave Output Terminal).  The output current is twice the set master current while the output voltage is the same as the master setting value.

To operate in the parallel mode:

1.	Set the TRACKING MODE switches on the front panel to PAR.

2.	Turn on the POWER switch.

3.	Turn the SLAVE Current  and Voltage Adjustment  Knobs clockwise to maximum.

4.	Set the master output voltage and current as described in section 3.2.

Figure 5     Parallel Connection

## 3.5	Fixed 5V/5A Output

This is the standard 5V/5A power output for use by TTL logic circuits.  When the load exceeds 5 amperes, the red OVERLOAD LED lights up.  The following figures show the parallel-output connection method and its characteristic curve.

Figure 6    Fixed 5V/5A Output Connection

## 3.6	Serial Operation

Serial connection is used to connect two or more power supplies in series to obtain a higher voltage ratings (maximum 240V for this power supply).

Figure 7 shows the serial connection between two 6303D power supplies at series tracking.  One unit is set at 60V/2A (MASTER setting is 30V/2A) while the other is at 40V/3A (MASTER setting is 20V/3A).  Figure 8 shows the serial connection voltage/current output versus load changes.

| Figure 7 Serial Connection Diagram | Figure 8 Serial Voltage/Current Output vs. Load Change |  |  |  |
| --- | --- | --- | --- | --- |

## 3.7	Parallel Operation

Parallel connection is used to connect two or more power supplies of the same model  in parallel to obtain a higher current output (maximum  24A for this power supply).

Figure 9 shows the parallel connection between two 6303D power supplies at parallel tracking.  One unit is set at 30V/5A (MASTER setting is 30V/2.5A) while the other is at 20V/6A (MASTER setting is 20V/3A).  Figure 10 shows the parallel connection voltage/current output versus load changes.

| Figure 9 Parallel Connection Diagram | Figure 10 Parallel Voltage/Current Output vs. Load Change |  |
| --- | --- | --- |

NOTE:	The difference between the voltage settings for two or more power supplies of the same model connected in parallel should not exceed 15V.  The minimum voltage setting should not be less than 10V.  If it is necessary to be lower than 10V, the voltage difference should be less than 2V.  The closer to 0V, the smaller the difference required.

# 4.	Maintenance

## 4.1	Changing the Fuse

The fuse is located inside the input power fuse holder (refer to Figure 2).   You need to change the fuse when:

·	the fuse is blown out

·	you change the voltage rating

In any case, replace the fuse with one of the same rating.  Refer to Table 1 for the type of fuse used for each model and voltage.

NOTE:	Unplug the power cord before you change the fuse.

## 4.2	Changing the Voltage

To change the voltage, follow these steps:

1.	Use a flathead screwdriver to detach the Input Power Fuse Holder and Input Voltage Selector  unit (refer to Figure 2).

2.	Turn the unit so that the desired input voltage is positioned above the  mark.

4.	Check that the fuse to conform to this new voltage rating.   Refer to Table 1 below for the correct fuse rating.

5.	Replace the Input Power Fuse Holder and Input Voltage Selector unit.

Table 1	Power Supply Weight , Dimensions and Fuse Specification

| Model | Weight (Approx.) |  |  | Dimensions W x H x D (mm) |  |  | Fuse Time-Delay Type 6 x 30 mm |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Net | Gross | Machine |  | Package | 100V |  | 120V | 220V | 240V |
| 6302A/AR/D/DR | 9.0 KG | 10.5 KG | 230x160x324 |  | 364x265x422 | 4A |  | 4A | 2A | 2A |
| 6303A/AR/D/DR | 9.0 KG | 10.5 KG | 230x160x324 |  | 364x265x422 | 5A |  | 5A | 2.5A | 2.5A |
| 6306A/AR/D/DR | 13 KG | 14.4 KG | 230x160x366 |  | 364x265x464 | 8A |  | 8A | 4A | 4A |
| 6603A/AR/D/DR | 13 KG | 14.4 KG | 230x160x366 |  | 364x265x464 | 8A |  | 8A | 4A | 4A |

# 5.	Specifications

| Model Specification | 6302A 6302D | 6303A 6303D | 6306A 6306D | 6603A 6603D |  | 6302AR 6302DR | 6303AR 6303DR | 6306AR 6306DR | 6603AR 6603DR |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Output Voltage | 0-30V x 2 | 0-30V x 2 | 0-30V x 2 | 0-60V x 2 |  | 0-30V x 2 | 0-30V x 2 | 0-30V x 2 | 0-60V x 2 |
| Output Current | 0-2A x 2 | 0-3A x 2 | 0-6A x 2 | 0-3A x 2 |  | 0-2A x 2 | 0-3A x 2 | 0-6A x 2 | 0-3A x 2 |
| CONSTANT VOLTAGE CHARACTERISTICS |  |  |  |  |  |  |  |  |  |
| Load Regulation | + 0.01% + 2 mV |  |  |  | + 0.01% + 2 mV |  |  |  |  |
| Line Regulation | + 0.01% + 2 mV |  |  |  | + 0.01% + 2 mV |  |  |  |  |
| Ripple & Noise (< 200W) | < 0.5mVrms |  |  |  | < 1mVrms |  |  |  |  |
| Ripple & Noise (> 200W) | < 1mVrms |  |  |  | < 1mVrms |  |  |  |  |
| CONSTANT CURRENT CHARACTERISTICS |  |  |  |  |  |  |  |  |  |
| Load Regulation (<200W) | < 10mA |  |  |  | < 10mA |  |  |  |  |
| Load Regulation (>200W) | < 15mA |  |  |  | < 15mA |  |  |  |  |
| Line Regulation | + 0.01% + 2mA |  |  |  | + 0.01% + 2mA |  |  |  |  |
| Ripple & Noise (< 200W) | < 1mArms |  |  |  | < 1mArms |  |  |  |  |
| Ripple & Noise (> 200W) | < 3mArms |  |  |  | < 3mArms |  |  |  |  |
| DISPLAY ACCURACY |  |  |  |  |  |  |  |  |  |
| Analog Display (A) Digital Display (D) | Full-scale 3% < 0.1% + 2d |  |  |  |  |  |  |  |  |
| PROGRAMMING SPEED |  |  |  |  |  |  |  |  |  |
| Rise Time (No Load) (Load) | < 100mS < 200mS (<6A), < 500mS (<10A), < 1S（> 10A) |  |  |  |  |  |  |  |  |
| Fall Time (No Load) (Load) | < 2.5S < 250mS |  |  |  |  |  |  |  |  |
| OUTPUT IMPEDANCE | < 2m + 2 µH |  |  |  |  |  |  |  |  |
| RECOVERY TIME | < 100 µS to within 0.1% of set voltage (50% to 100% load change) |  |  |  |  |  |  |  |  |
| FUNCTIONS |  |  |  |  |  |  |  |  |  |
| Series Connection | Different models can be connected in series (< 240V) |  |  |  |  |  |  |  |  |
| Parallel Connection | Same models can be connected in parallel (< 24A) |  |  |  |  |  |  |  |  |
| Master - Slave | N A |  |  |  | Yes |  |  |  |  |
| REMOTE CONTROL |  |  |  |  |  |  |  |  |  |
| Voltage Programmable By External DC Voltage By External Resistance | NA |  |  |  | 0 to 10V, Control 0 to voltage setting 0 to 10K, Control 0 to maximum voltage |  |  |  |  |
| Current Programmable By External DC Voltage By External Resistance | NA |  |  |  | 0 to 10V, Control 0 to current setting 0 to 10K, Control 0 to maximum current |  |  |  |  |
| TRACKING OPERATION |  |  |  |  |  |  |  |  |  |
| Tracking Error | + 0.2% + 10mV |  |  |  |  |  |  |  |  |
| Series Regulation | + 0.3% + 10mV |  |  |  |  |  |  |  |  |
| 5V FIXED OPERATION |  |  |  |  |  |  |  |  |  |
| Load Regulation | < 10mV |  |  |  |  |  |  |  |  |
| Line Requlation | < 5mV |  |  |  |  |  |  |  |  |
| Ripple & Noise | < 2mVrms |  |  |  |  |  |  |  |  |
| Voltage Accuracy | 5V + 0.25V |  |  |  |  |  |  |  |  |
| Output Current | 5A |  |  |  |  |  |  |  |  |

| Model Specification | 6302A 6302D | 6303A 6303D | 6306A 6306D | 6603A 6603D |  | 6302AR 6302DR | 6303AR 6303DR | 6306AR 6306DR | 6603AR 6603DR |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Output Voltage | 0-30V x 2 | 0-30V x 2 | 0-30V x 2 | 0-60V x 2 |  | 0-30V x 2 | 0-30V x 2 | 0-30V x 2 | 0-60V x 2 |
| Output Current | 0-2A x 2 | 0-3A x 2 | 0-6A x 2 | 0-3A x 2 |  | 0-2A x 2 | 0-3A x 2 | 0-6A x 2 | 0-3A x 2 |
| VOLTAGE/CURRENT READBACK AVAILABLE |  |  |  |  |  |  |  |  |  |
| Voltage Readback | NA |  |  |  | 0 to 10V for full-scale voltage in Master |  |  |  |  |
| Current Readback | NA |  |  |  | 0 to 0.3V for full-scale current in Master |  |  |  |  |
| OPERATION MODE |  |  |  |  |  |  |  |  |  |
| Independent | Two independent outputs and 5V fixed output Output from 0 to rating voltage and 0 to rating ampere |  |  |  |  |  |  |  |  |
| Serial | Output from 0 to + rating voltage at rating ampere Output from 0 to double rating voltage at rating ampere |  |  |  |  |  |  |  |  |
| Parallel | Output from 0 to double rating current at voltage rating |  |  |  |  |  |  |  |  |
| POWER SOURCE | ACV 100/120/220/240 + 10%, 50/60 Hz |  |  |  |  |  |  |  |  |
| ACCESSORIES | ACS-002 X 2 |  |  |  | ACS-002 X 2, 9-pin D-type connector x 2 |  |  |  |  |
| DIMENSIONS W X H X D (mm) | 230 X 160 X 324 |  | 230 X 160 X 366 |  |  | 230 X 160 X 324 |  | 230 X 160 X 366 |  |